## Slide 1: Irradiance Meter
Alex Polyanskiy
Montana State University
Bozeman, MT

## Slide 2: Project Overview
Design a prototype irradiance meter
  Research
  Select irradiance sensor and design circuit
  Incorporate an inclinometer
  Incorporate temperature sensor for ambient temperature and solar panel temperature
  Build prototype
  Test prototype

## Slide 3: Solar Irradiance
Solar Irradiance is radiant flux (power) per unit area () received from the sun

## Slide 4: Measuring Solar Irradiance
A pyranometer measure global irradiance (direct + diffused)
Irradiance is the integral of the spectrum over all wavelengths
Thermopile | Photodiode | Reference Cell
Flat Spectral Response Highest Accuracy Most Expensive Not used in handheld meters | Spectrally Selective, not what a panel sees Least Accurate Least Expensive Requires angle correction | Closely matches PV panel spectral response Accurate to PV Panels Cost depends on the Cell Calibration Sees what a solar panel sees
4

## Slide 5: Spectral Sensitivity
5

## Slide 6: Types of panels available
Monocrystalline (Mono-Si)
  20% or greater
  Expensive
  Space- efficient
Polycrystalline (P-Si)
  around 15%
  Process to make is cheaper and simpler
Amorphous
  7-10%
  Best if space is not an issue
  Degrade faster
Cadmium Telluride (CdTe)
  14-18%
  Disposal and recycling is costly and dangerous
Copper Indium Gallium Selenide (CIGS)
  up to 22.4%
  Many manufacture have gone bankrupt
Gallium Arsenide (GaAs)
  up to 28.9%
  Most expensive
  Used on spacecrafts
Thin-Film
6

## Slide 7: Spectral responsivity of different solar cells
Irradiance is the integral of the spectrum over all wavelengths
7

## Slide 8: Calculating Irradiance
Proportional to short-circuit current (IEC 60891):
With temperature compensation (IEC 60891):
Standard Test Conditions (IEC 61829)
  module temperature: 25 °C
  in-plane irradiance: 1 000 W/m2
  spectral distribution: AM 1,5 (global)
8

## Slide 9: Mismatch Factor (IEC 60904-7)
Geff as ref spectrum is the effective irradiance of the test specimen spectrum;
Gmeas is the irradiance as measured by the reference device;
MM is the spectral mismatch factor.
Eref(λ) is the irradiance per unit bandwidth at a particular wavelength λ,
of the reference spectral irradiance distribution;
Emeas(λ) is the irradiance per unit bandwidth at a particular wavelength λ, of the spectral irradiance distribution of the incoming light at the time of measurement;
S­­ref(λ) is the spectral response of the reference PV device;
Ssample(λ) is the spectral response of the test PV device.
In the case, that absolute spectra and absolute spectral responses are used for the analysis:
9

## Slide 10: Functional Overview
10

## Slide 11: First Design
11

## Slide 12: First Design
Microcontroller
Solar Cell
Inclinometer
Temperature Sensor
LCD
12

## Slide 13: Devices on the Market
Solmetric SolSensor
(photodiode)
Seaward Solar Survey200R
(reference cell)
13

## Slide 14: Results
14
Chart: Irradiance Measurement Comparison
| July 22, 1:00pm | July 23, 5:00pm | July 29, 11:00am (Wind) | July 29, 5:30pm | July 30, 10:00am | August 1, 10:30am (wind)
Solsensor (Photodiode) | 865 | 75 | 755 | 475 | 155 | 555
Seaward (Reference Cell) | 886 | 0 | 754 | 450 | 162 | 539
First Design (Reference Cell) | 853 | 65 | 756 | 480 | 153 | 565
S - low | 855 | 55 | 745 | 465 | 135 | 525
Sea - low | 880 | 0 | 750 | 446 | 154 | 522
I - low | 849 | 50 | 733 | 472 | 134 | 538
Inv. Value | 935

## Slide 15: Measurement Comparison (average of 6.8 W/m2)
Chart: Irradiance Difference Between Design and Solmetric
| July 22, 1:00pm | July 23, 5:00pm | July 29, 11:00am (Wind) | July 29, 5:30pm | July 30, 10:00am | August 1, 10:30am (wind)
Irradiance Difference | -9 | -7.5 | -5.5 | 6 | 1.5 | -11.5
15

## Slide 16: Second Design
16

## Slide 17: Second Design
LDC
Solar Cell
Bluetooth
Buttons
Mounting Holes
Thermocouple
17

## Slide 18: Second Design
Front View
Back View
18

## Slide 19: Ways to improve design
Add a lowpass filter
4-wire contact for shunt resistor
Test Cell at standard conditions
Compare to a secondary standard
Use electrical components with lower error
19

## Slide 20: Additional Tasks
Solar Panel Research
DC Combiner Box Research
Solar Simulator Research
Calibration Techniques
20

## Slide 21: Internship Take-aways
Design Review, what to consider, how to select parts
PCB Design
Error Analysis
Working with Computer Engineering and Mechanical Engineering
Corporate Environment
Meetings
21